| SOFTWARE DEVELOPMENT PLAN |  |
| --- | --- |
| Project number / Title | MT_ONE_0023 / miRFlow: A Novel Screening Tool for Endometrial Cancer |

| Version Control |  |  |  |  |
| --- | --- | --- | --- | --- |
| Version | Description |  | Date (dd Mmm yyyy) | Author |
| v0.01 | First draft for circulation |  | 08/03/2026 | Samin Yasar |
| v1.01 |  |  |  |  |
| APPROVAL |  |  |  |  |
| Name | Organisation | Position |  | Signature / Date |
| Dr Sylvain Ladame | Imperial College | Principal Investigator |  |  |

CONTENTS

References	4

References (providing input)	5

1.	INTRODUCTION	6

1.1	Scope	6

1.2	Terminology/Abbreviations/Acronyms / Definitions	6

2.	SOFTWARE OVERVIEW	6

3.	RESPONSIBILITIES & AUTHORISED SIGNATORIES	6

4.	APPLICABLE STANDARDS	7

5.	SOFTWARE DEVELOPMENT TOOLS	7

5.1	Software Development tools	7

5.2	Software configuration Management	7

5.3	Obsolescence management	8

5.4	Software development rules and standards	8

6.	SOFTWARE DEVELOPMENT PROCESS	8

6.1	Software development overview and design reviews	8

6.2	Software integration process	9

6.3	Software Verification	10

6.4	Risk Management	10

7.	Software Change Management	11

7.1.1	Version control	11

8.	Defect Management	11

# References

| Ref/ID | Title |
| --- | --- |
| IEC 62304:2006 +AMD1:2015 | Medical device software — Software life cycle processes |
| EN 62304: 2006+A1:2015 | Medical device software — Software life cycle processes |
| FDA Guidance | Guidance for the Content of Premarket Submissions for Software Contained in Medical Devices |

# References (providing input)

| Ref/ID | Title |
| --- | --- |
| QSP0005 | Risk Management |
| QSP0006 | Design and Development Control |
| QSP0019 | Software Design and Development Control |
| FM1002 | Msk MEC xxx, Key project responsibilities |
| FM1010 | Msk MEC xxx, Intended use, Customer and regulatory requirements |
| FM1013 | Msk MEC xxx, Design Specification_y.yy |
| FM1015 | Msk MEC xxx, Design Plan_y.yy |
| FM0051 | Design Reviews |
| N/a | Meeting minutes |

# INTRODUCTION

## Scope

This document presents the software development plan for the miRFlow project, which includes software developed as a standalone analytical application with a planned graphical user interface (GUI) for processing and analysing flow cytometry data generated by the miRFlow diagnostic assay. It covers the design, development, and release of the software. This document provides a detailed plan for the software components. The top-level plan for software development is integrated into the overall Design Plan for the project.

## Terminology/Abbreviations/Acronyms / Definitions

| Terminology/ Abbreviations/ Acronyms/ Definitions | Meaning |
| --- | --- |
| AWS | Amazon Web Services |
| PHI | Protected Health Information |
| BLE | Bluetooth Low Energy |
| VCS | Version Control System |
| CPU | Central Processing Unit |
| OS | Operating System |
| PC | Personal Computer |
| USB | Universal Serial Bus |
| RCR | Risk Control Requirement (cf risk analysis) |
| SRS | Software Requirement Specification |
| FC | Flow Cytometry |
| FCS | Flow Cytometry Standard File |
| MMB | Magnetic Micro-Bead (Size Signal) |
| DG | Dragon Green-Bead (Fluorescence Signal) |
| GMM | Gaussian Mixture Model |
| ROI | Region of Interest |
| ML | Machine Learning |

# SOFTWARE OVERVIEW

The miRFlow software is being developed as a standalone analytical application designed to process and interpret flow cytometry (FC) data generated by the novel miRFlow diagnostic assay which detects microRNA (miRNA) biomarkers associated with endometrial cancer. The software converts raw FC signal outputs into structured analytical results and clinically interpretable metrics.

The software operates by analysing fluorescence signals produced by magnetic microbeads (MMBs) and Dragon Green (DG) beads within the assay. These signals reflect the formation of miRNA-mediated sandwich complexes, which are indicative of the presence of specific miRNA targets in patient samples. By processing the fluorescence and size signals recorded during FC acquisition, the software identifies and quantifies bead-binding states and generates probabilistic metrics representing the likelihood of miRNA biomarker presence.

The miRFlow software is currently being developed to run on a standard computing platform (e.g., a personal computer running a compatible operating system with Python support). The system processes Flow Cytometry Standard (.fcs) files exported from FC instruments used in laboratory settings. The software does not directly control the FC instrument but instead analyses exported datasets produced by the instrument.

The system architecture is modular, consisting of several backend processing modules that communicate sequentially to transform raw data into structured outputs. These modules include data ingestion and visualisation, signal standardisation, bead classification, probabilistic clustering, and diagnostic scoring. Each module performs a defined stage of the analysis pipeline and produces intermediate outputs that are passed to subsequent modules.

At the current stage of development, the primary focus is on implementing the backend analytical pipeline responsible for signal processing and probabilistic classification of bead-binding states. Development of a graphical user interface (GUI) is planned for a later stage of the project and will allow clinicians or laboratory personnel to upload FC datasets, execute the analysis pipeline, visualise results, and export diagnostic reports.

In addition to the FC-based analytical pipeline, the system architecture includes a planned secondary analysis branch intended to integrate confocal microscopy-derived imaging data. This component will support future machine learning (ML)-based event weighting, enabling multi-modal analysis that combines FC signals with imaging-derived binding metrics to improve diagnostic robustness. This functionality is currently under development and will be implemented at the later stages of the project.

The final system will generate structured analytical outputs including labelled visualisations, summary data files, and diagnostic ratios derived from the quantified miRNA sandwich signals. These outputs will support interpretation of the assay results and may be presented through the planned GUI interface for clinical review.

# RESPONSIBILITIES & AUTHORISED SIGNATORIES

The Project Management and lead software development responsibilities are described in Msk MEC xxx, Key project responsibilities (FM1002). In addition, the following responsibilities are given for software development:

| Activity | Name of the person Responsible for activity | Position | Change (if applicable) |
| --- | --- | --- | --- |
| Configuration tools management | Samin Yasar | Lead software engineer |  |
| Setting up development tools | Samin Yasar | Lead software engineer |  |
| Software specifications approval | Dr Sylvain Ladame | PI |  |

Authorised signatories for the Design and Development records, including software development records, are listed in 1.3, MD_NPD_001 Authorised signatories, in section 1 of the Design History File (DHF). All design and development inputs and all design and development outputs are reviewed and approved by the PI, to ensure the requirements of QSP0006 are fulfilled.

# APPLICABLE STANDARDS

A full list of applicable standards and guidances for the xxx, including software development standards and guidances, are listed in Msk MEC xxx, Applicable standards_y.yy (FM1005).

# SOFTWARE DEVELOPMENT TOOLS

## Software Development tools (Pre-requisites)

The following tools are used during the software development:

| Item | Tool | Details | Version used |
| --- | --- | --- | --- |
| Workstation | Personal Computer (PC) – Windows-based workstation | Minimum requirements: Multi-core CPU, ≥8 GB RAM, ≥10 GB available storage. (Used for development, data processing, and visualisation of flow cytometry outputs) | Windows 10 / Windows 11 |
| Requirements management and documentation | Microsoft Word (Microsoft) | Used for drafting software documentation including SDP, SA, and project reports. | Microsoft Office 365 |
| Software design | Microsoft PowerPoint / Microsoft Visio (Microsoft) | Used for system architecture diagrams, workflow diagrams, and module interaction visualisation for the miRFlow software pipeline. | Microsoft Office 365 |
| Coding and automated tests | Python (Python Software Foundation), Visual Studio Code (Microsoft) | Python used for development of all analytical modules including data ingestion, signal transformation, gating, clustering, and scoring. Coding standards follow standard Python conventions (PEP 8). | Python 3.x, VS Code (latest stable) |
| Software configuration management (SCM) | Git / GitHub (GitHub Inc.) | Version control system used to track code revisions, maintain development history, and ensure traceability of software changes. | Git 2.x |
| Supporting items | Python scientific libraries: NumPy, Pandas, SciPy, scikit-learn, Matplotlib, fcsparser, flowutils | Used for numerical computation, data handling, flow cytometry file parsing, logicle transformation, statistical modelling, and data visualisation. | Latest stable versions (to be documented at final release) |

## Software configuration Management

Software Configuration Management (SCM) for the miRFlow software project is implemented using Git, with code repositories hosted on GitHub. Git provides version control capabilities that allow tracking of software changes, management of development history, and traceability of modifications made throughout the software lifecycle.

The SCM system provides a command-line interface and integrates with development environments used during implementation. It enables structured management of software artefacts including source code, configuration files, and documentation.

The repository maintains a master development branch that represents the current stable version of the software. Additional branches may be created when major updates or new features are being developed. These branches allow isolated development and testing before merging updates back into the main branch.

A simplified branching model is currently used during development. As the project progresses, additional branch management strategies may be introduced if required.
[Placeholder: optional branch workflow diagram may be added here if needed.]

All relevant software configuration items are maintained under configuration control within the SCM system before they are incorporated into the released version of the software. These configuration items include:

- Source code for analytical modules

- Configuration files (e.g., JSON gating thresholds and analysis parameters)

- Supporting scripts and data processing tools

- Documentation related to software development and architecture

Each configuration item is uniquely identifiable within the SCM system and version controlled to ensure traceability of changes across development iterations.

Where Software of Unknown Provenance (SOUP) is used, such as third-party Python libraries and external software dependencies, the software name, vendor, and version will be documented. These dependencies may include scientific computing libraries used for data processing, statistical modelling, and visualisation.

Documentation relating to SOUP components and other supporting development items will be maintained within the project’s software documentation directory.

[Placeholder: final list of SOUP dependencies and versions to be documented prior to final software release.]

## Obsolescence management

Software development tools used in the miRFlow project will be managed to minimise risks associated with software obsolescence while maintaining development stability and reproducibility.

During active development, stable versions of the programming environment, libraries, and supporting tools will be used to ensure consistent behaviour of the analytical pipeline. Where updates to development tools or third-party software dependencies become available, these updates will be evaluated before adoption to determine whether they introduce improvements, security updates, or compatibility issues with the existing software modules.

Where appropriate, development tools may be updated to newer stable versions if the update provides clear benefits and does not negatively affect the reproducibility or functionality of the software. Any such updates will be recorded through the software configuration management (SCM) system to ensure traceability of the development environment.

In cases where updates could introduce incompatibility or instability, the currently validated version of the tool will continue to be used for the remainder of the development phase. This approach ensures that the behaviour of the analysis pipeline remains consistent throughout testing and validation activities.

All software development tools and third-party dependencies used in the miRFlow system will be documented, including their respective versions, prior to final software release.

[Placeholder: final list of software development tools and dependency versions to be documented at project completion.]

## Software development rules and standards

The miRFlow software is developed following a set of structured development practices to ensure code readability, maintainability, and reproducibility of the analytical pipeline. These practices define conventions for software modelling, coding standards, data handling, and documentation.

- Software Modelling and Architecture Documentation

High-level software architecture and workflow interactions are documented using system diagrams and workflow visualisations. These diagrams describe the modular structure of the software pipeline and the interactions between individual processing modules, including data ingestion, signal transformation, gating, clustering, and scoring components.

Where appropriate, modelling diagrams may be created using standard diagramming tools to represent the logical structure of the system and the flow of data between modules. These diagrams support software documentation and assist with communication of system design.

[Placeholder: additional modelling diagrams or architecture documentation may be added during later development stages.]

- Coding Standards

All software modules for the miRFlow project are implemented in Python. Coding practices follow widely accepted Python development conventions to promote readability and maintainability of the codebase.

Key coding practices include:

- Adherence to PEP 8 Python style guidelines

- Use of clear and descriptive variable and function names

- Modular design with clearly separated functional components

- Use of inline comments and docstrings to document functions and processing steps

- Consistent formatting of code to improve readability and collaboration

Where possible, functions are designed to perform a single logical task, supporting modularity and ease of debugging.

- Data Handling and File Structure Standards

Standardised file structures are used to organise input data, intermediate outputs, and final analysis results. The software processes Flow Cytometry Standard (.fcs) files as input and produces structured outputs including visualisation files, JSON configuration outputs, and aggregated summary tables.

Intermediate configuration parameters such as gating boundaries and bead thresholds are stored in structured configuration files (e.g., JSON format) to ensure reproducibility and traceability of the analysis pipeline.

- Documentation and Traceability

Software documentation is maintained alongside the source code and includes descriptions of the software architecture, processing modules, and configuration files used in the analysis pipeline. Version control is used to maintain traceability of code updates and development iterations.

Where third-party libraries or external software dependencies are used, these components are documented as Software of Unknown Provenance (SOUP) within the project documentation.

[Placeholder: additional documentation standards or testing guidelines may be introduced during later development stages.]

# SOFTWARE DEVELOPMENT PROCESS

## Software development overview and design reviews

The software development for the miRFlow project follows the organisation’s Software Design and Development control process (QSP 0019). This document defines the structured approach used to manage software design, development, and documentation activities throughout the project lifecycle.

The development of the miRFlow analytical software follows an iterative development methodology with agile characteristics. This approach was selected because the project involves research-stage algorithm development and experimental assay optimisation, where incremental development and continuous evaluation of intermediate results are necessary. Modules of the software pipeline are implemented, tested, and refined progressively as new experimental datasets become available and the system architecture evolves.

This methodology allows individual components of the pipeline (such as data ingestion, signal transformation, gating, clustering, and scoring modules) to be developed independently and integrated progressively into the overall analysis framework. Iterative refinement also supports validation of algorithmic performance using experimental flow cytometry data generated during the project.

The software development activities form part of the wider design and development process for the miRFlow diagnostic system, which aims to analyse flow cytometry outputs from the miRFlow assay and generate clinically interpretable metrics for microRNA detection. All documentation generated during the development process is stored within the relevant sections of the Design History File (DHF) as indicated in the lifecycle table below.

The lifecycle of software development is split into the following phases:

| Software Life cycle phase | Design phase (see QSP0005) | Generated records & files | Design Review point (see also the Design Plan) |
| --- | --- | --- | --- |
| Software Concept and Planning | Design and Development Planning (DHF section 6) | Software development Plan (FM3001) | Design Review @ Project Review gate |
| Software Requirements Analysis | Design and Development inputs (DHF section 5) | Software Requirement Specification (FM3006) Risk analysis (FM0021) Software architecture | Design Review @ Project Review gate |
| Software Design | Project Implementation (DHF section 8) | Design Phase specifications (FM1016) to cover: Software detailed design Software integration detailed design if required | Design Review at Design Phases closure for each phase |
| Software Implementa-tion | Project Implementation (DHF section 8) | Design History Log (Software) (FM0041) Software and its configuration files Attachments to the Design Phase Reports: Automated Test results sheets Code review meeting minutes | Review at Design Phases / Work Package closure |
| Software Integration* | Project Implementation (DHF section 8) | Software integration Test plan report | Work Package Specification review |
| Software Verification | Design and Development verification (DHF section 10) | Software verification test plan (FM3010), test protocols and test report Software system testing verification reports | Design Review after Verification summary report |
| Software Validation | Design and Development validation (DHF section 12) | Master validation plan (FM0025), Record of design validation form (FM1026) | Design Review @ Critical Review gate** |
| Software Release | Design and Development Validation (DHF section 11) | Software release form FM3020, including the software and its configuration files, user guides, Installation procedures, admin procedure and reviewing all test results to verify that no critical bugs are left open and that all required tests have been conducted. | Design Review @ Critical Review gate** |

*For classes B and C

** The following activities shall be considered during this phase:

•Software verification shall be completed, and the results evaluated before the Software is released.

•Validation activities shall be completed prior to software release for review in critical gate review.

•All known residual anomalies shall be documented.

•All known residual anomalies shall be evaluated to ensure that they do not contribute to an unacceptable risk

•The version of the software product that is being released shall be documented.

•The procedure and environment used to create the released software shall be documented.

•All activities and tasks defined in the Software Development Plan are complete along with all the associated documentation.

•The Software Product and configuration items shall be archived. Associated documentation shall be archived in line with MD_QP03, for at least a period determined as the longer of:

•The life-time of the device; or

•A time specified by relevant regulatory requirements.

The software development phases indicated above are integrated into the Design plan (FM1015), which prescribes the traceability from user requirements to validation, as well as the design phases and gate reviews (preliminary, project and critical review) during the overall design and development process. The documents are filed in the corresponding section of the DHF.

## Software integration process

During the coding and integration stages of the miRFlow software development, integration activities are performed progressively as individual software modules reach functional completion. The purpose of integration testing is to ensure that independently developed components of the analysis pipeline interact correctly and produce consistent outputs when combined within the overall system architecture.

The software architecture follows a modular pipeline structure, where individual processing units perform specific stages of the analytical workflow. These modules include data ingestion and visualisation, signal standardisation, gating boundary extraction, probabilistic clustering, and final scoring calculations. Integration testing is performed as these modules are connected sequentially to form the full analysis pipeline.

Integration testing focuses on verifying that:

- Data outputs generated by one module are correctly interpreted by downstream modules

- Configuration files and intermediate data structures are correctly transferred between modules

- The integrated pipeline produces consistent outputs when processing representative flow cytometry datasets

Integration tests are performed using experimental and control Flow Cytometry Standard (.fcs) datasets, ensuring that the end-to-end pipeline behaves as expected when analysing real assay data.

Prior to integrating major changes into the primary development branch of the software repository, the proposed updates are reviewed to ensure that the changes are compatible with the existing architecture and do not disrupt previously implemented functionality. Version control tools are used to manage these integration updates and maintain traceability of modifications.

Where new Software of Unknown Provenance (SOUP) components or third-party dependencies are introduced, these components are reviewed to ensure compatibility with the existing software environment before integration into the analysis pipeline.

Results from integration testing activities will be documented and maintained as part of the project documentation.

[Placeholder: detailed integration test report and formal integration test records to be completed during the final stages of the project.]

## Software Verification

Software verification activities for the miRFlow software will follow the verification process defined for the overall project. The purpose of software verification is to confirm that the implemented software modules perform according to the defined software requirements and that risk control measures identified during risk analysis are correctly implemented.

Verification activities will demonstrate that the software implementation conforms to the specifications defined in the Software Requirement Specification (SRS) and the associated risk control requirements. Verification activities and outcomes will be documented to maintain traceability between software requirements, verification protocols, and verification results.

Each software requirement and risk mitigation associated with the software will be linked to one or more verification activities to ensure full traceability from requirement → verification protocol → verification report.

At the current stage of the project, the verification framework is being defined while core modules of the analytical pipeline are still under active development. Formal verification testing will be conducted once the software pipeline reaches functional completion.

[Placeholder: final verification protocols and verification reports will be completed during the final stages of the project.]

- Software Unit Verification

Software unit verification focuses on confirming the correct behaviour of individual software modules within the analysis pipeline. These modules include components responsible for data ingestion, signal transformation, gating boundary generation, bead classification, probabilistic clustering, and scoring calculations.

Unit verification activities may include:

- Code reviews to verify correctness of implementation

- Execution of test datasets to confirm expected module behaviour

- Inspection of intermediate outputs generated by individual modules

- Verification that configuration files and parameters are correctly applied

Acceptance criteria for unit verification include confirmation that the module produces expected outputs when tested with representative datasets and that no critical processing errors are observed.

[Placeholder: detailed unit test procedures and results will be documented in the final verification report.]

- Software Integration Verification

Software integration verification confirms that individual modules operate correctly when combined into the full analysis pipeline. Integration verification ensures that data outputs from upstream modules are correctly interpreted by downstream modules and that configuration files and intermediate data structures are passed correctly between components.

Integration testing will be performed using representative Flow Cytometry Standard (.fcs) datasets to confirm correct end-to-end behaviour of the pipeline.

Integration verification activities are closely related to Software System Verification, where the complete analytical workflow is tested as a unified system.

[Placeholder: integration verification procedures and results will be recorded in the integration test documentation.]

- Software System Verification

Software system verification evaluates the behaviour of the full miRFlow analytical pipeline when processing flow cytometry datasets. These tests confirm that the integrated system performs the intended signal processing, event classification, and diagnostic scoring functions.

System verification will include evaluation of algorithm performance using representative experimental datasets and comparison against defined evaluation metrics for clustering stability and signal separation.

[Placeholder: detailed verification protocols and system-level testing results will be defined in the final verification plan.]

- Acceptance Criteria for Verification Testing

The acceptance criteria for verification testing will include the following conditions:

- All verification protocols identified in the verification plan have been executed using a controlled version of the software build.

- Where necessary, a code freeze will be applied to the software build during verification testing to ensure that only critical fixes are introduced and that additional verification is performed following any such changes.

- Issues identified during verification testing will be reviewed and addressed where necessary, with corrective actions documented.

- Any unresolved issues will be reviewed and assessed to confirm that they do not introduce unacceptable risks or compromise the intended functionality of the software.

[Placeholder: final verification results and confirmation of acceptance criteria will be documented upon completion of verification testing.]

## Risk Management

Risk management for the miRFlow software follows the risk management process defined for the overall project. Software-related risks are identified and evaluated as part of the broader system risk analysis for the miRFlow diagnostic assay.

Potential software-related hazards are assessed to determine whether failures in the analytical pipeline could lead to incorrect interpretation of assay results or incorrect system outputs. Identified risks are analysed and appropriate mitigation measures are incorporated into the software design and implementation.

Where risks are mitigated through software functionality or algorithm design, the corresponding control measures are documented within the Software Requirement Specification (SRS). These requirements ensure that risk control measures are implemented within the software architecture and analytical pipeline.

All software requirements, including those generated as part of risk control activities, are documented within the Technical Requirement Specification. This documentation provides traceability between identified risks, implemented control measures, and the verification activities used to confirm that the requirements have been satisfied.

[Placeholder: detailed software risk analysis documentation to be completed as part of the final project risk assessment.]

# Software Change Management

During software design and development, changes to the miRFlow software are tracked using the Software Configuration Management (SCM) system. All updates to the codebase are recorded through version control commits, ensuring traceability of modifications made throughout the development process.

Changes to the software are periodically reviewed during development to ensure that updates are consistent with the intended system architecture and do not negatively affect existing functionality. Version control history allows tracking of the origin, purpose, and impact of software changes.

Where significant design or functionality changes are introduced, their potential impact on the software system and analytical pipeline will be reviewed to ensure that the changes do not introduce unintended behaviour or compromise previously implemented features.

Formal change management processes may be applied during later stages of the project once the software reaches a stable release stage.

[Placeholder: formal change control procedures may be introduced if the software transitions toward a production or clinical deployment stage.]

### Version control

During development, revisions to the software are tracked using the commit history within the SCM repository. Each commit represents a recorded change to the codebase, allowing traceability of software evolution throughout the project.

When a stable build of the software is identified, the corresponding SCM commit revision may be linked to a defined software version number.

Version numbers for release candidates follow the semantic versioning format:

MAJOR.MINOR.PATCH

The version number components are updated according to the following rules:

- MAJOR version – incremented when incompatible or major structural changes are introduced

- MINOR version – incremented when new functionality is added in a backward-compatible manner

- PATCH version – incremented when backward-compatible bug fixes or minor corrections are applied

Additional labels may be used for pre-release or testing versions (e.g., alpha, beta, or release candidate) as extensions to the standard version format.

Where the software consists of multiple independent modules within the analysis pipeline, these modules may maintain internal version tracking where appropriate. However, the overall system release will be associated with a single version identifier representing the integrated software build.

# Defect Management

Software defects or anomalies identified during development, testing, or integration activities will be documented and tracked to ensure transparency and traceability throughout the development lifecycle.

When a defect is identified, the issue will be recorded with sufficient detail to allow investigation and resolution. The recorded information will include a description of the issue, the conditions under which it occurred, and the potential impact on software functionality or system outputs.

Where possible, defects will be corrected through updates to the software codebase, with the changes tracked through the Software Configuration Management (SCM) system. Verification or re-testing may be performed following the implementation of corrective changes to confirm that the issue has been resolved and that no new issues have been introduced.

All unresolved anomalies will be recorded in the design history documentation, indicating the problem, the potential impact on device performance, any mitigation strategies or workarounds, and an assessment of the potential risk to device safety or effectiveness.

[Placeholder: a formal defect log or issue tracking register may be maintained during later development stages if required.]